Ao-bergamo DM calibration SOP

1. Assume that the microscope is aligned, both the 2p and measurement paths (which should be as parallel as possible).
  1. You can run both DM measurement laser and 2p ex laser at the same time to verify this.
  2. Verify that the DM is in the Fourier plane of the RGG via the white DM proxy. Interference pattern should not change when the scanner is turned on.
  3. Verify the objective pupil plane is conjugate the DM – ‘test deformable mirror” in shwfs program should not change image @ 38mm below the PMT arm.
2. Assumes that the wavefront sensor image is centered and very nearly flat (centroids should be close to the calibrated locations)
  1. The microlens array is glued in, so this should not need to be updated. It was set using the Zygo interferometer in the Jet optics room.
3. If you are happy with mechanical / optical alignment of everything, it’s time to calibrate the forward transfer function of the DM
  1. Start shwfs.
  2. Click ‘matlab mmap interface’ and ‘test deformable mirror’ checkboxes.
  3. The wavefront should begin to wiggle.
  4. Click ‘record data’. The “data size” counter should begin incrementing.
  5. Allow > 5000 data frames (combination of both the DM command signal and wavefront) to be recorded. Give it about half an hour..
  6. Uncheck ‘record data’
  7. Click ‘write data’ to save rundata/centroids.mat
  8. Open up matlab.
  9. Run ‘Cforward_fit.m’
  10. It will give diagnostics, including how many centroids are observed. This needs to be >= 530 for reliable operation.
    1. If it’s < 530, adjust the alignment on the WFS arm such that the image is centered and not clipping. This can be tricky, as the WFS arm is using 1” optics that are nearly 100% filled. But, with care, you can get 535-540 visible centroids.
    2. If you need to adjust the optics, go back to the beginning above for recalibration of the forward transfer function of the DM.
  11. Other diagnostics include the quality of the linear prediction of the DM control signals from the wavefront sensor. The relative error in both plots should be low, < 5%. If it’s not, redo the procedure.
  12. If the calibration looks good, save it and proceed.
  13. Close shwfs program.
4. Next step is to calibrate the flat or system-correction wavefront. This is more complicated, as it relies on feedback from ScanImage.
  1. Make sure both the DM control laptop and ScanImage computer are both plugged into the same ethernet switch, and both have IP addresses on the ethernet network.
    1. If wifi is on for linux, it won’t automatically pull in a IP address from DHCP on the ethernet port (provided through a USB-C dongle, presently). You’ll need to locate the eth inteterface via ‘sudo ip a’, then run ‘dhclient’ on that name.
  2. Open Scanimage in Matlab 2019a on the windows computer.
  3. Open Matlab on the Linux computer and locate geneopt.m
  4. Load the ‘geneopt_dmnet’ configuration into ScanImage
  5. Open the user functions dialog in SI, and uncheck everything.
  6. Place a beads sample of the desired color / wavelength under the microscope (w the large block of aluminum as a pedestal, e.g.), and use epi to locate the focal plane.
  7. Open up shwfs, check the DM and WFS are working.
  8. Enable ‘matlab mmap interface’ in shwfs
  9. Check and uncheck ‘zero DM command’
  10. Switch to 2p mode, regular SI style, and image the beads with the DM command zeroed. It will be blurry – the DM by default is not flat.
  11. Check the IP address of the Linux computer. Open imageStats_reset on the windows computer, and make sure that the IP address is correct there.
  12. Open imageStatsFcn on the windows computer, and make sure the correct channel is selected for feedback (1= red, 2= green)
  13. Make sure both red and green pmts are checked for display.
  14. Make sure both imageStats_reset and imageStatsFcn are enabled (and nothing else) in the User functions dialog of SI.
  15. On the Linux computer, run ‘geneopt’. It should say ‘ok go’ when ready.
  16. Press ‘Focus’ in Scanimage. This will begin acquiring data, which is sent to the Linux computer. The linux computer will apply perturbations to the DM & measure their effect on the image brightness. This is used as selective criteria in genetic optimization to improve the commands signals to the DM, while also recording the resulting wavefront.
  17. The whole process takes several hours. When done, change the last lines in geneopt.m to save the resulting data in a file with the date and wavelength (e.g.).
  18. Using this file name (you have to edit the file again, sorry I’m lazy & this works), run Cgeneopt_analyze.m on the linux computer. This will select & average top solutions from the genetic optimization run, and save it to a calibration file in ao-bergamo/data
  19. Now you need Zernike coefficients for this ‘flat’ or system correction wavefront. Run Czernike.m on the Linux computer.
  20. Close shwfs
  21. You’ll need to edit ao-bergamo/src/dm-control.cpp (sorry again!) to include the new calibrations into online control. The relevant lines are at the top of the file and should be self-explanatory.
  22. Run ‘make clean’ then ‘make’ in ~/ao-bergamo to rebuild the executable.
  23. Restart shwfs. The new calibration should be available from the drop-down menu & it should deform the DM when ‘closed loop control of DM’ is enabled. If not, examine stdout for information on which variables didn’t load properly from the matlab files.