AO-Bergamo two-photon microscope SOP

Purpose: This documents the standard procedures for operating the AO-Bergamo in 2C.357 for imaging live cranial windowed mice.

1. Preparation
  1. At least 45 minutes before imaging, turn the laser (Coherent Discovery TPC) from standby to enable using the key switch on the laser power supply (below the table as you enter the room)
    1. Verify using the Coherent Discovery GUI app (ScanImage computer; icon on the desktop)
    2. If the laser doesn’t start warming up, or isn’t connected, you will need to turn it off (switch on the back of the black laser power supply), wait a few minutes, and turn it back on.
      1. It will make a loud noise, the red lights on the front of the power supply will blink, and the GUI app should work.
      2. The chiller (below the laser power supply) and MRU (hidden above the table) should never be turned off.
2. System checks – make sure everything is working before you anesthetize your subject.
  1. Laser: make sure the laser is connected via the Coherent Discovery GUI and operating with a reasonable power
    1. > 1.8W 900-1000nm, > 1.5W 1000-1100nm, > 1.3w 1100-1300nm
  2. AO: On the laptop, open a terminal (if one is not opened) and type shwfs (‘shack-hartmann wavefront sensor’). This will open a window where you can see the wavefront, as represented as the lenslet centroids.
    1. Make sure the mirror moves when you check ‘test deformable mirror’. Uncheck when done.
      1. If it does not, go to ‘Debugging the deformable mirror’, below.
    2. Check and un-check ‘zero DM command’. This should flatten (to some approximation) the mirror.
    3. Check ‘closed loop control of DM’, and select a wavefront calibration. Verify that the wavefront changes then is stable.
    4. Close shwfs.
      1. Only one program open at a time! (It conflicts with the leveling optics / camera)
  3. Anesthesia: check there is oxygen in the tank outside the door, and isoflurane in the vaporizer
    1. The vaporizer has a sight glass with two white lines for upper and lower limits; refill if it’s too low.
      1. There are more iso bottles by the west entrance to the vivarium.
  4. Headbar orientation: if the headbar is in the front, the aluminum tube that holds the mouse needs to be moved to the back position.
    1. Likewise, if the headbar is to the back of the head, the tube should be in the forward position.
    2. You can move it by unscrewing the two Phillips screws at the bottom of the tube – long handle screwdriver should be in the rig – and switching the holes that the screws go into.
  5. Heat: if the mouse will be fully anestetized, turn on the aluminum tube heater. The controller should be set for 38C. Make sure it’s heating up.
    1. If the mouse is awake, do not turn the heater on, it will only irritate the rodent.
  6. Nose cone: Take the nose cone off the surgical rig beside the sink, store the scavenge hose adapter and set-screw, and plug the inlet / scavenge tubes into the small and large polyethylene tubes by the microscope.
    1. Small polyethylene tube = isoflurane; large = scavenge.
    2. Plug the other ends of the PE tubes into the vaporizer, via the angle adapter, and into the wall vacuum by sliding one tube inside the other.
    3. Mount the nose-cone on the leveling goiniometer using the thorlabs thumb-screw.
      1. This jig is adjstable for different mice.
    4. In this way, you can route anesthesia + oxygen to the microscope & nose-cone.
3. Procedure
  1. Anesthetize the subject using the induction box.
    1. I like 4% iso, Filip used 5%. 0.7-0.8 lpm oxygen flow. Be sure to check the outlet valves of the vaporizer to route the anesthetic correctly.